CIN NO : U33204MH2022PTC387499              CERTIFICATE OF ANALYSIS

| Product Name | 5 ml LT COC Barrel |  |  |
| --- | --- | --- | --- |
| Batch No | P370011221 |  |  |
| Effective date: | 12/2021 | Review before: | 11/2026 |
| Shelf Life | 60 Months |  | Page no: 1 of 3 |

| Sr. No. | Tests | Results | Specifications |
| --- | --- | --- | --- |
| 1 | Description | Complies | Transparent plastic barrel of 5 ml. free from any visual defect and should be well clean and having smooth surface, free from cracks or pin holes |
| 2 | Leakage Test | Complies | There should not be any signs of leakage from any container observed after 24 hours |
| 3 | Appearance of solution | Complies | Solution should be clear and colorless and should be practically free from foreign particles. |
| 4 | Acidity or Alkalinity | 0.1 ml | Not more than 0.3ml of 0.01M sodium hydroxide or 0.01M hydrochloric acid should be require to change the color of indicator |
| 5 | Absorbance | 0.20 | The absorbance at 230 nm to 360 nm should not be more than 0.40 |
| 6 | Reducing substances | Complies | The difference between the titration volumes should not be more than 3.0 ml |
| 7 | Transparency | Complies | The cloudiness of the suspension should be perceptible when viewed through the container and compared with a similar container filled with water |
| 8 | Nonvolatile Residue | Complies | The difference between the amount obtained from the sample and the blank should not exceed 15 mg |
| 9 | Residue on Ignition | Complies | The difference between the amount of residue on ignition obtained from the sample and the blank should not exceed 5 mg. |

|  | Prepared By | Reviewed By | Checked By | Approved By |
| --- | --- | --- | --- | --- |
| Name |  |  |  |  |
| Signature |  |  |  |  |
| Date |  |  |  |  |
| Section | Quality Control | Quality Control | Quality Control | Quality Assurance |

210, Samruddhi Commercial Complex, Mind Space, Off Link Road, Malad (W). Mumbai – 400064. India.

www.addishealthcare.com

| Product Name | 5 ml LT COC Barrel |  |  |
| --- | --- | --- | --- |
| Batch No | P370011221 |  |  |
| Effective date: | 12/2021 | Review before: | 11/2026 |
| Shelf Life | 60 Months |  | Page no: 2 of 3 |

| 10 | Heavy Metals | 0.6 ppm | Not more than 1 ppm |
| --- | --- | --- | --- |
| 11 | Buffering Capacity | Complies | The difference between the two volumes should not be greater than 10 ml, and if acid was required for either the sample or blank and alkali for the other, the total of the two volumes required should be greater than 10 ml |
| 12 | Total Length (l) (L2) | 76.80 mm | 76.70 mm ± 0.20 nm |
| 13 | Length of barrel (l1) (L1) | 66.55 mm | 66.50 mm ± 0.20 nm |
| 14 | Outer diameter (d1) (D1) | 14.40 mm | 14.40 mm ± 0.10 nm |
| 15 | Inner diameter (d2) (D2) | 11.75 mm | 11.75 mm ± 0.10 nm |
| 16 | Flange diameter (d3) (D3) | 23.15 mm | 23.10 mm ± 0.10 nm |
| 17 | Flange cut diameter (d4) (D4) | 19.40 mm | 19.40 mm ± 0.10 nm |
| 18 | Flange thickness (h1) (t1) | 2.50 mm | 2.50 mm± 0.15 nm |
| 19 | Over flow capacity | 7.2 ml | Not less than 7.0 ml |
| 20 | Safety test | Complies | During the observation period none of the animals treated with the extract shows a significantly greater biological reactivity than the animals treated with the blank. |
| 21 | Bacterial Endotoxin Test | Less than 0.25 EU/barrel | Not more than 0.25 EU/barrel |
| 22 | Particulate Matter |  |  |
|  | Particle Size: ≥ 10µm | 300 | Not more than 600 per barrel |
|  | Particle size: ≥ 25µm | 20 | Not more than 60 per barrel |
| 23 | Sterility Test | Sterile | Should be sterile |

|  | Prepared By | Reviewed By | Checked By | Approved By |
| --- | --- | --- | --- | --- |
| Name |  |  |  |  |
| Signature |  |  |  |  |
| Date |  |  |  |  |
| Section | Quality Control | Quality Control | Quality Control | Quality Assurance |

210, Samruddhi Commercial Complex, Mind Space, Off Link Road, Malad (W). Mumbai – 400064. India.

www.addishealthcare.com

| Product Name | 5 ml LT COC Barrel |  |  |
| --- | --- | --- | --- |
| Batch No | P370011221 |  |  |
| Effective date: | 12/2021 | Effective date: | 12/2021 |
| Shelf Life | 60 Months |  | Page no: 3 of 3 |

| 24 | Ethylene Oxide | Complies | Not more than 10 ppm |
| --- | --- | --- | --- |
| 25 | Residue of Silicone Oil | 1.7 mg/barrel | Not more than 2 mg/barrel |
| 26 | Gliding Force | 5.5 N | 10 Newton |
| 27 | Siliconization powder test | Complies | Should be proper, homogenized sticking of the powder in inner surface of barrel. |
| Conclusion: The sample complies with the above mention test specifications |  |  |  |

210, Samruddhi Commercial Complex, Mind Space, Off Link Road, Malad (W). Mumbai – 400064. India.

www.addishealthcare.com

|  | Prepared By | Reviewed By | Checked By | Approved By |
| --- | --- | --- | --- | --- |
| Name |  |  |  |  |
| Signature |  |  |  |  |
| Date |  |  |  |  |
| Section | Quality Control | Quality Control | Quality Control | Quality Assurance |